MEASUREMENT PROTOCOL 	Date

Hospital	:

System / Detector	:	Tube Above/Below:

Procedure	:

Physician (initials)	:

1. PROTECTIVE EQUIPMENT

|  |  |  | mm Pb |
| --- | --- | --- | --- |
| Lead Apron |  |  |  |
| Thyroid Collar |  |  |  |
| Eyewear |  |  |  |
| Lead Gloves |  |  |  |
| Table Curtain |  |  |  |
| Patient Shield |  |  |  |
| Ceiling |  |  |  |
| Floor |  |  |  |

2. STAFF POSITION

| Number of persons in the room / position1 |  |
| --- | --- |
| MDs : | A |
|  | B |
|  | C |
| Techn.: | A |
|  | B |
|  | C |
| Nurses: | A |
|  | B |
|  | C |
| Others: | A |
|  | B |
| Access 2 (R / F / P / IJV) |  |

3. GENERAL PARAMETERS

| FSD3 |  |  |
| --- | --- | --- |
| FDD3 |  |  |
| FOV |  |  |
| ZOOM |  |  |
| Biplane (Y/N) |  |  |
| DAP present (Y/N)? | Tube 1 | Tube 2 |

4A. FLUOROSCOPY

| Projection4 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| kV |  |  |  |  |  |
| mA |  |  |  |  |  |
| Pulses/s |  |  |  |  |  |
| Pulse duration ms |  |  |  |  |  |
| Filtration mm Cu |  |  |  |  |  |
| Mode |  |  |  |  |  |

B. CINE

| Projection 4 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| kV |  |  |  |  |  |
| mA |  |  |  |  |  |
| Frames/s |  |  |  |  |  |
| No of Frames |  |  |  |  |  |
| Pulse duration ms |  |  |  |  |  |
| Filtration mmCu |  |  |  |  |  |
| Mode |  |  |  |  |  |

________________________

1 Define position as shown on picture e.g. MDs: A 4, B 5…

2 Radial (R) / Femoral (F) / Percutaneous (P) / Internal Jugular Vein (IJV) access

3 FSD = Focus-Skin Distance, FDD = Focus-Detector Distance

4 Projections (PA, AP, LAO, RAO…) and degrees (o): e.g. LAO 30o/ CAUD 20o

5. DAP VALUES (units: ……..)

| FLUORO : |  |  | Fluoroscopy Time (….…) : |  |
| --- | --- | --- | --- | --- |
| CINE : |  |  | Total Frames Acquired : |  |
| TOTAL : |  |  |  |  |

6. OTHER REMARKS

Body part irradiated	: 	   Physician’s Height	:

Patient size 		   Complexity of procedure

(child / S / M / L)	: 	   (Low / Medium / High)	:

Physician’s experience 		   Physician’s

(Low / Medium / High)	: 	   Dominant Hand (R / L/ Both)	:

comments:

7. TLDs – MEASURING POINTS - RESULTS

TLD type 	:

Wrist

Ring

Eyes

Leg

Number of TLDs	:

TLD positions 	  Pellet No	 Measured Dose (mSv)

R	 L 	 R             L

- Ring finger

- Wrist

Lead Apron

Leg

Middle   R/L side	 Middle     R/L side

- Eyes

Outside surface of finger when the tube is above the table;

Inside surface of finger when the tube is below the table. (TLDs are stuck on MD’s fingers)

Outside surface of wrist when the tube is above the table;

Inside surface of wrist when the tube is below the table.

Part of the leg not covered by lead apron

Two measuring points as shown on picture, under the lead glasses:  	1  if the tube is on the right side of the MD

2  if the tube is on the left side of the MD.